## Slide 1
Measure the length of the screen diagonal in cm
Notes: ScreenSizeMeasurement.png

## Slide 2
Measure the distance from the eyes of the subject to the screen
Distance in cm
Notes: ScreenPositionInstruction.png

## Slide 3
Measure the required position of the fiducial with the VTG Calibrator
Procedure:
Place the VTG Calibrator in the screen corner where the fiducials are to be generated.
Read of the vertical and horizontal distances from the centre of the fiducial edge of display area.
Please see the LabBench VTG manual if further information is required on the calibration procedure.
Notes: PositionInstruction.png

## Slide 4
Update the LabBench DISPLAY configuration in the LabBench Designer program
1
2
3
4
Notes: Update instruction

## Slide 5
Start the test to generate a fiducial and use the VTG calibrator to confirm the fiducial is in the correct position
Notes: FiducialInstruction.png

## Slide 6
Mount the LabBench VTG device on the screen and start this test.
This test will generate a series of visual stimuli with fiducials and will verify that triggers are correctly generated by the LabBench VTG.
Each stimuli will consist of displaying a number. Each time you see a number press any button on the LabBench PAD.
Notes: DetectorInstruction.png

## Slide 7
1
Notes: S1.png

## Slide 8
2
Notes: S2.png

## Slide 9
3
Notes: S3.png

## Slide 10
4
Notes: S4.png